## Metadata
Batch Type | Assay
Batch Record Version | 2.0.3
Date | 2026-03-26
Run by | AJ
Machine | Zoo QX600
Batch ID | AB142
Analysis Software Version | 2
Input uL | 5 | * Amount of extraction included in each reaction. Can be overriden for single wells in the "Volume Overrides" sheet.
Reaction uL | 20
Amplification Method | Bio-Rad OneStep RT-ddPCR Advanced Kit for Probes
Amplification Method Lot ID | 64667524
Quantification Type | Direct Quantification
Assay Method | GTM GT-Digital Flu;SC2;RSV Wastewater Surveillance Panel v1.0 for the Bio-Rad QX600 (100735)
Assay Method Lot ID | 22184
Batch Comment | None
Target Fluorophore | FAM
Target | Influenza Virus A (M)
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | None
Control 1 Macromolecule | None
Control 1 Well | H4
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | Oligomer
Control 2 Macromolecule | RNA
Control 2 Well | G4
Control 2 Predicted Copies Per uL Reaction | 100
Target Fluorophore | HEX
Target | Respiratory Syncitial Virus, Bovine
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | None
Control 1 Macromolecule | None
Control 1 Well | H4
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | Oligomer
Control 2 Macromolecule | RNA
Control 2 Well | G4
Control 2 Predicted Copies Per uL Reaction | 600
Target Fluorophore | Cy5
Target | Respiratory Syncitial Virus, Human (G)
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | None
Control 1 Macromolecule | None
Control 1 Well | F4
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | Oligomer
Control 2 Macromolecule | RNA
Control 2 Well | G4
Control 2 Predicted Copies Per uL Reaction | 600
Target Fluorophore | Cy5.5
Target | Influenza Virus B (NEP/NS1)
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | None
Control 1 Macromolecule | None
Control 1 Well | F4
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | Oligomer
Control 2 Macromolecule | RNA
Control 2 Well | G4
Control 2 Predicted Copies Per uL Reaction | 600
Target Fluorophore | ROX
Target | SARS-CoV-2 (SC2)
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | None
Control 1 Macromolecule | None
Control 1 Well | F4
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | Oligomer
Control 2 Macromolecule | RNA
Control 2 Well | G4
Control 2 Predicted Copies Per uL Reaction | 600

## Sample IDs
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | A01-1519 | C01-1265 | F01-1008 | N02-1003
B | A01-1520 | C01-1266 | D01-1016 | M04-1003
C | A01-1521 | C01-1267 | K01-0478 | J01-1017
D | A01-1522 | C01-1268 | P01-1012 | R02-1018
E | A01-1523 | G01-1005 | O01-1005 | Q01-1017
F | A01-1524 | H01-1024 | N03-1001 | NTC
G | A01-1525 | U01-0170 | L01-1010 | PC
H | C01-1264 | W01-1005 | M04-1002 | NTC

## Comments
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Volume Overrides
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Storage Locations
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | NAB164 L1 | NAB165 A1 | B1 | C1
B | L2 | A2 | B2 | C2
C | L3 | A3 | B3 | C3
D | L4 | A4 | B4 | C4
E | L5 | A5 | B5 | C5
F | L6 | A6 | B6
G | L7 | A7 | B7
H | L8 | A8 | B8

## Run Setup
ddPCR Formulation Worksheet
Enter or confirm any values that appear in the blue cells.
Computed values will appear in the orange cells and should not be changed manually.
Separate assays/reaction: | 1 | <- Enter the number of separate assays (e.g., primer/probe mixes) that are included in each reaction. This number is usually 1, and will automatically adjust the volume of water to add to the mastermix.
Mastermix in each reaction mix (µL): | 16.5 | These numbers are almost never changed! Each reaction is prepared with 10% extra volume to allow for transfer inefficiencies during assay run.
Sample in each reaction mix (µL): | 5.5
Total volume per reaction mix (µL): | 22
Total volume per assay (µL): | 20
Reaction Conditions
Step | Temp (°C) | Time (min) | Cycles
Reverse Transcription: | 50 | 60 | 1
RT Denaturation: | 95 | 10 | 1
Strand Denaturation: | 94 | 0.5 | 40
Anneal & Extend: | 55 | 1
Finish: | 98 | 10 | 1
Thermocycler Settings
Ramp: | 2°C/sec
Lid: | 104°C
Reaction vol: | 40 µL
Total reactions in run: | 32 | Do not change this number manually! Enter the sample IDs for the current run on the Plate worksheet, including all negative and positive controls. This will auto-update the total number of reactions shown here. and affect the computed component volumes below.
Mastermix overrun: | 1.1 | Compensation for liquid transfer errors. Do not change this without consulting the PI!
Mastermix Composition (µL)
Water: | 232
Supermix: | 194 | Vortex 30 seconds just before adding
Reverse Transcriptase (RT): | 77.5 | Vortex 30 seconds just before adding
DTT (300 mM): | 38.8 | Vortex 30 seconds just before adding
Each assay (20X): | 38.8 | Vortex 30 seconds just before adding
Total mastermix volume: | 581 | Vortex final mastermix 30 seconds before use

## Lists
Machine IDs | Targets | Categories | Target Macromolecules | Control Templates | Control Types | Fluorophores | Amplification Methods | Assay Types | Assays | QxManager Versions | Hosts
Zoo QX200 | SARS-CoV-2 (SC2) | Pathogen | None | None | PCR Negative | FAM | Bio-Rad OneStep RT-ddPCR Advanced Kit for Probes | Direct Quantification (dPCR; ddPCR) | GTM GT-Digital Flu;SC2;RSV Wastewater Surveillance Panel v1.0 for the Bio-Rad QX600 (100735) | 2 | NA
Zoo QX600 | SARS-CoV-2 (N1) | Gene or Genomic Region | DNA | Oligomer | PCR Positive | HEX | Bio-Rad ddPCR Supermix for Probes | Indirect Quantification (real-time PCR) | GTM GT-Digital MeV WT Wastewater Surveillance Assay v1.0 for the QX200/QX600 (100840) | 1.2 | Human
Zoo Qiacuity 1 | SARS-CoV-2 (N2) | Internal Control | RNA | SARS-CoV-2 | Other | Cy5 | ----------------------------- | Rare Mutation Detection | GTM dPCR NVO;Mpox Assay Kit v1.0 for the QX200-BETA (100458) |  | Avian
Other | Human (RNase P) | Other | Protein | LV-SARS_N |  | Cy5.5 | QIAcuity OneStep Advanced Probe Kit | Other | GTM dPCR NVO;Mpox Assay Kit v1.0 for the QX200-BETA (100458) plus GT-Digitial Mpox Clade I OPG210 ROX Add-On for the QX600 (100865) |  | Swine
| Human Immunodeficiency Virus 1 |  | Other | Influenza A |  | ROX | ----------------------------- |  | GTM dPCR Norovirus GI;GII Assay Kit for the Bio-Rad QX200 (100468) |  | Multiple
| Human Immunodeficiency Virus 2 |  |  | Influenza B |  | ATTO 590 | Other |  | ----------------------------- |  | Other
| Influenza Virus A (M) |  |  | PMMoV |  | FAM/HEX |  |  | Bio-Rad Influenza A (M), B (NEP/NS1), Human (RNAse P) Triplex (dEXD84844642)
| Influenza Virus B (NEP/NS1) |  |  | Respiratory Syncitial Virus, Bovine |  | Cy5/Cy5.5 |  |  | Bio-Rad SARS-CoV-2 (HVdel69-70) (dMDS944624402)
| Measles Virus, Wild-Type (N) |  |  | Other |  |  |  |  | Bio-Rad SARS-CoV-2 (N1, N2) Duplex (dEXD15704188)
| Mpox Virus Clade I (F3L) | [7 empty] | Bio-Rad SARS-CoV-2 (N1, N2), BRSV (N) Triplex (dEXD41434415)
| Mpox Virus Clade II (F3L) | [7 empty] | Bio-Rad SARS-CoV-2 (N1, N2), Human (RNAse P) Triplex (dEXD28563542)
| Non-Variola Orthopoxviruses | [7 empty] | Bio-Rad SARS-CoV-2 (N1) (dEXD54243734)
| Norovirus GI, Human (ORF1_2) | [7 empty] | Bio-Rad SARS-CoV-2 (N2) (dEXD28513507)
| Norovirus GII, Human (ORF1_2) | [7 empty] | Bio-Rad Variant SARS-CoV-2 Multiplex (dEXD15841942)
| Pepper Mild Mottle Virus | [7 empty] | -----------------------------
| Respiratory Syncitial Virus, Bovine | [7 empty] | GTM GT-Digital infA;infB Assay Kit v1.1 for the QIAcuity-BETA (100519-1.1)
| Respiratory Syncitial Virus, Human (G) | [7 empty] | GTM GT-Digital MeV WT M Gene WW Surveillance Assay v1.0 for the Qiacuity-BETA (101035)
| Other | [7 empty] | -----------------------------
Custom Design