## Metadata
Batch Type | Assay
Batch Record Version | 2.0.0
Date | 1899-12-31
Run by | TD
Machine | Zoo QX200
Analysis Software Version | 2.0
Batch ID | AB###
Input uL | 5 | * Amount of extraction included in each reaction. Can be overriden for single wells in the "Volume Overrides" sheet.
Reaction uL | 20
Amplification Method | Bio-Rad OneStep RT-ddPCR Advanced Kit for Probes
Amplification Method Lot ID | xxxxxxx
Quantification Method | ddPCR
Quantification Type | Direct Quantification
Method | SARS-CoV-2 (N1, N2), Human (RNAse P) Triplex (dEXD28563542)
Method Lot ID | xxxxxxx
Batch Comment | None
Target Fluorophore | FAM
Target | SARS-CoV-2
Target Genetic Locus | N1:SARS
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | NA
Control 1 Macromolecule | NA
Control 1 Well | XX
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | LV-SARS_N
Control 2 Macromolecule | RNA
Control 2 Well | XX
Control 2 Predicted Copies Per uL Reaction | 100
Target Fluorophore | FAM/HEX
Target | SARS-CoV-2
Target Genetic Locus | N2:SARS
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | PCR Negative
Control 1 Template | NA
Control 1 Macromolecule | NA
Control 1 Well | XX
Control 1 Predicted Copies Per uL Reaction | 0
Control 2 Type | PCR Positive
Control 2 Template | LV-SARS_N
Control 2 Macromolecule | RNA
Control 2 Well | XX
Control 2 Predicted Copies Per uL Reaction | 600
Target Fluorophore | HEX
Target | Homo sapiens
Target Genetic Locus | RNase P
Target Macromolecule | RNA
Target Predicted Copies Per uL Reaction | 0
Control 1 Type | NA
Control 1 Template | NA
Control 1 Macromolecule | NA
Control 1 Well
Control 1 Predicted Copies Per uL Reaction
Control 2 Type | NA
Control 2 Template | NA
Control 2 Macromolecule | NA
Control 2 Well
Control 2 Predicted Copies Per uL Reaction

## Sample IDs
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Extraction IDs
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Comments
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Volume Overrides
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A
B
C
D
E
F
G
H

## Run Setup
ddPCR Formulation Worksheet
Enter or confirm any values that appear in the blue cells.
Computed values will appear in the orange cells and should not be changed manually.
Separate assays/reaction: | 1 | <- Enter the number of separate assays (e.g., primer/probe mixes) that are included in each reaction. This number is usually 1, and will automatically adjust the volume of water to add to the mastermix.
Mastermix in each reaction mix (µL): | 16.5 | These numbers are almost never changed! Each reaction is prepared with 10% extra volume to allow for transfer inefficiencies during assay run.
Sample in each reaction mix (µL): | 5.5
Total volume per reaction mix (µL): | 22
Total volume per assay (µL): | 20
Reaction Conditions
Step | Temp (°C) | Time (min) | Cycles
Reverse Transcription: | 50 | 60 | 1
RT Denaturation: | 95 | 10 | 1
Strand Denaturation: | 94 | 0.5 | 40
Anneal & Extend: | 55 | 1
Finish: | 98 | 10 | 1
Thermocycler Settings
Ramp: | 2°C/sec
Lid: | 104°C
Reaction vol: | 40 µL
Total reactions in run: | 0 | Do not change this number manually! Enter the sample IDs for the current run on the Plate worksheet, including all negative and positive controls. This will auto-update the total number of reactions shown here. and affect the computed component volumes below.
Mastermix overrun: | 1.1 | Compensation for liquid transfer errors. Do not change this without consulting the PI!
Mastermix Composition (µL)
Water: | 0
Supermix: | 0 | Vortex 30 seconds just before adding
Reverse Transcriptase (RT): | 0 | Vortex 30 seconds just before adding
DTT (300 mM): | 0 | Vortex 30 seconds just before adding
Each assay (20X): | 0 | Vortex 30 seconds just before adding
Total mastermix volume: | 0 | Vortex final mastermix 30 seconds before use

## Lists
Machine IDs | Targets | Categories | Loci | Target Macromolecules | Control Templates | Control Types | Fluorophores | Amplification Methods | Assay Quantification | Assay Types | Assays | QxManager Versions | Hosts
Zoo QX200 | NA | NA | NA | NA | NA | NA | FAM | Bio-Rad OneStep RT-ddPCR Advanced Kit for Probes | ddPCR | Direct Quantification | Adenovirus (dEXD92003507) | 1.2 | NA
Zoo QX600 | Adenovirus | Pathogen | 16S | DNA | crAssphage | PCR Negative | HEX | Bio-Rad ddPCR Supermix for Probes | RT-PCR | Rare Mutation Detection | Enterococcus (23S) (dEXD14924415) | 2.0 | Human
Other | Antibiotic Resistance Gene (AbR) | Gene or Genomic Region | 23S | RNA | GBlock | PCR Positive | Cy5 | Other | Other | Indirect Quantification | HF183 (16S) (dEXD79194188) |  | Avian
| crAssphage | Internal Control | C3b | Protein | Human RNase P | Other | Cy5.5 |  |  | Other | Influenza A (M), B (NEP/NS1), Human (RNAse P) Triplex (dEXD84844642) |  | Swine
| Enterococcus | Other | (crAssphage) | Other | LV-SARS_N |  | ROX |  |  |  | Legionella (16S) (dEXD53463961) |  | Multiple
| Enterovirus |  | E |  | PMMoV |  | ATTO 590 |  |  |  | MPXV Clade I (C3b) (dEXD77548788) |  | Other
| Escherichia coli, Shiga toxin-producing (STEC) |  | G |  | RSV |  | FAM/HEX |  |  |  | MPXV Clade II (TNFR) (dEXD51818561)
| Homo sapiens |  | H |  | SARS-CoV-2 |  | Cy5/Cy5.5 |  |  |  | Norovirus (POL) (dEXD27733734)
| Human Immunodeficiency Virus 1 (HIV-1) |  | H1 |  | Other | [5 empty] | RSV A/B (G) (dEXD77482599)
| Human Immunodeficiency Virus 2 (HIV-2) |  | M:FLUA | [7 empty] | SARS-CoV-2 (HVdel69-70) (dMDS944624402)
| Influenza Virus A (FluA) |  | N | [7 empty] | SARS-CoV-2 (N1, N2) Duplex (dEXD15704188)
| Influenza Virus B (FluB) |  | N1:SARS | [7 empty] | SARS-CoV-2 (N1, N2), BRSV (N) Triplex (dEXD41434415)
| Legionella |  | N2:SARS | [7 empty] | SARS-CoV-2 (N1, N2), Human (RNAse P) Triplex (dEXD28563542)
| Mpox Virus I (MPXV) |  | NEP/NS1 | [7 empty] | SARS-CoV-2 (N1) (dEXD54243734)
| Mpox Virus II (MPXV) |  | (PMMoV) | [7 empty] | SARS-CoV-2 (N2) (dEXD28513507)
| Norovirus |  | POL | [7 empty] | Variant SARS-CoV-2 Multiplex (dEXD15841942)
| Pepper Mild Mottle Virus (PMMoV) |  | RNase P | [7 empty] | Custom Design
| Respiratory Syncitial Virus, Bovine (BRSV) |  | S
| Respiratory Syncitial Virus, Human (RSV) |  | S:E484K
| Rhinovirus |  | S:E484K WT
| Salmonella enterica subsp. enterica serovar Enteritidis |  | S:HVdel69-70
| Salmonella enterica subsp. enterica serovar Javiana |  | S:HVdel69-70 WT
| Salmonella enterica subsp. enterica serovar Newport |  | S:N501Y
| Salmonella enterica subsp. enterica serovar Typhimurium |  | S:N501Y WT
| SARS-CoV-2 |  | TNFR
| Other |  | Other